**PRODUCT CONFORMITY CERTIFICATE**

This is to certify that the

***Thermo-FID TOC Analyser***

***PT (Portable Thermo-FID), FE (Thermo-FID model Field Housing), ES (Thermo-FID model 19” Rack Mount), TG (Thermo-FID Desk Top), MK (Thermo-FID model Close Coupled Probe On Stack), KA (Thermo-FID model Cassette Mount)***

manufactured by:

***SK-Elektronik GmbH***

Benzstraße 23 -25

51381 Leverkusen

Germany

has been assessed by Sira Certification Service

and for the conditions stated on this certificate complies with:

**MCERTS Performance standards and test procedures for continuous emission**

**monitoring systems (CEMS) and transportable-CEMs (T-CEMS), October 2020**

**EN 15267-2:2007, EN15267-3:2007,**

& QAL 1 as defined in EN 14181: 2014

Certification ranges :

TOC 0 to 15 mg/m3

0 to 30 mg/m3

0 to 2000 mg/m3

|  |  |  |
| --- | --- | --- |
|  |  |  |
| Project number: | 80057891 |
| Certificate number: | Sira MC050062/05 |
| Initial certification: | 03 October 2005 |
| This certificate issued: | 02 October 2020 | Andrew Young |
| Renewal date: | 02 October 2025 | Environmental Team Manager |

**Certificate contents**

[Approved site application 2](#_Toc377369165)

[Basis of certification 2](#_Toc377369166)

[Product certified 3](#_Toc377369167)

[Certified performance 4](#_Toc377369168)

[Description 7](#_Toc377369169)

[General notes 7](#_Toc377369170)

# Approved site application

*Any potential user should ensure, in consultation with the manufacturer, that the monitoring system is suitable for the intended application. For general guidance on monitoring techniques refer to the Environment Agency technical guidance on monitoring, available at* [*www.mcerts.net*](http://www.mcerts.net)

This instrument is considered suitable for use on waste incineration and large combustion plant applications. This CEMS has been proven suitable for its measuring task (parameter and composition of the flue gas) by use of the QAL 1 procedure specified in EN14181. The lowest certified range for each determinand shall not be more than 1.5 times the daily average emission limit value (ELV) for incineration plants, and not more than 2.5 times the ELV for other types of application.

# Basis of certification

This certification is based on the following test report(s) and on Sira’s assessment and ongoing surveillance of the product and the manufacturing process:

TUV Rheinland Report Number 936/806016 dated 26 February 1997


TUV Rheinland Report Number 936/806016/B dated 23 December 2003

TUV Rheinland Report Number 936/21219522/A dated 21 June 2013

TUV Rheinland Doc number 20150413\_936 dated 13 April 2015

# Product certified

The Thermo-FID TOC measuring system consists of the following parts:



|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **1. Sample Probe** | **2. Heated Filter** | **3. Heated Sample Line** | **4. Gas Conditioning** | **5. Analyser** |
| Gasy MKA | Gasy MKA | Gasy MKA | n/a | Thermo FID |

Note: This only applies to the TG, PT, FE and ES model. The MK consists only of one stack-mountable unit.

Allowable variations could include:

* A different brand or model of sampling system of the same type, provided that there is evidence the alternative system works with similar types of CEM.
* Additional manifolds and heated valves used to allow more than one analyser to share a sampling system.

This certificate applies to all instruments fitted with software version 5.00 onwards (serial number bearing the year ‘04’ onwards).

For EN 15267-3, this certificate applies to all instruments fitted with software version 5.31 onwards (serial number 1803412 onwards).

# Certified performance

The instrument was evaluated for use under the following conditions:

Ambient Temperature Range: +5°C to +45°C

Instrument IP rating: IP40 (ES, PT or TG); IP54 (FE or MK (IP65 Version available))

Note: For outdoor installations the analyser needs to be mounted into an IP65 environment. If the instrument is supplied with an enclosure, then the ambient temperature shall be monitored inside the enclosure to ensure that it stays within the above ambient temperature range.

Results are expressed as error % of certification range, unless otherwise stated.

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Test | Results expressed as % of the certification range | | | | Other results | MCERTS  specification |
|  | <0.5 | <1 | <2 | <5 |  |  |
| Response time |  |  |  |  |  |  |
| TOC (0 to 15 mg/m3) |  |  |  |  | 12s | <200s |
| TOC (0 to 30 mg/m3) |  |  |  |  | 13s | <200s |
| TOC (0 to 2000 mg/m3) |  |  |  |  | 14s | <200s |
| Repeatability standard deviation at zero point |  |  |  |  |  |  |
| TOC | 0.2 |  |  |  |  | <2.0% |
| Repeatability standard deviation at reference point |  |  |  |  |  |  |
| TOC | 0.2 |  |  |  |  | <2.0% |
| Lack-of-fit |  |  |  |  |  |  |
| TOC (0 to 15 mg/m3) | 0.2 |  |  |  |  | <2.0% |
| TOC (0 to 30 mg/m3) | 0.3 |  |  |  |  | <2.0% |
| TOC (0 to 2000 mg/m3) |  |  | 1.0 |  |  | <2.0% |
| Influence of ambient temperature zero point |  |  |  |  | Note 1 |  |
| TOC |  |  |  | 4.8 |  | <5.0% |
| Influence of ambient temperature reference point |  |  |  |  | Note 1 |  |
| TOC |  |  |  | -2.0 |  | <5.0% |
| Influence of sample gas flow for extractive CEMS |  |  |  |  |  |  |
| TOC | -0.1 |  |  |  |  | <2.0% |
| Influence of voltage variations 190 to 250V |  |  |  |  |  |  |
| TOC |  |  | -1.0 |  |  | <2.0% |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Test | Results expressed as % of the certification range | | | | Other results | MCERTS  specification |
|  | <0.5 | <1 | <2 | <5 |  |  |
| Influence of vibration (10 to 60Hz (±0.3mm), 60 to 150Hz at 19.6m/s2) |  | -0.7 |  |  |  | To be reported |
| Cross-sensitivity at zero with interferents: O2, H2O, CO, CO2, N2O, NO, NO2, NH3, SO2, HCl |  |  |  |  | Note 2 |  |
| TOC | 0.0 |  |  |  |  | <4.0% |
| Cross-sensitivity at reference with interferents: O2, H2O, CO, CO2, N2O, NO, NO2, NH3, SO2, HCl |  |  |  |  | Note 2 |  |
| TOC |  |  |  | 3.9 |  | <4.0% |
| Effect of oxygen for TOC CEMS |  |  | 1.3 |  |  | <2.0% |
| Response factors for TOC CEMS |  |  |  |  |  |  |
| Methane |  |  |  |  | 1.0 | 0.9 to 1.2 |
| Aliphatic hydrocarbons |  |  |  |  | 1.04 – 1.10 | 0.9 to 1.1 |
| Aromatic hydrocarbons |  |  |  |  | 0.84-1.08 | 0.8 to 1.1 |
| Dichloromethane |  |  |  |  | 1.09 | 0.75 to 1.15 |
| Aliphatic alcohols |  |  |  |  | 0.90-1.04 | 0.7 to 1.0 |
| Esters and keytones |  |  |  |  | 0.94-0.99 | 0.7 to 1.0 |
| Organic acids |  |  |  |  | 0.75-0.77 | 0.5 to 1.0 |
| Test gas mixture |  |  |  |  | <13.6% | <15% |
| Measurement uncertainty |  |  |  |  | Guidance - at least 25% below max permissible uncertainty | |
| TOC |  |  |  |  | 11.8% | <22.5% (<30%) |
| Calibration function (field) |  |  |  |  |  |  |
| TOC |  |  |  |  | 0.9904-0.9929 | >0.90 |
| Response time (field) |  |  |  |  |  |  |
| TOC |  |  |  |  | <170s | <200s |
| Lack of fit (field) |  |  |  |  |  |  |
| TOC |  |  | 1.8 |  |  | <2.0% |

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Test | Results expressed as % of the certification range | | | | Other results | MCERTS  specification |
|  | <0.5 | <1 | <2 | <5 |  |  |
| Maintenance interval |  |  |  |  | Four weeks | >8 days |
| Zero and Span drift requirement | The system is equipped with an automatic drift correction. A status signal is triggered when the zero and span points are outside the specified ranges.  The system is able to perform an automatic calibration with the supplied test gas. If the deviation is higher than 10%, the instrument triggers the status signal for maintenance. If the deviation is higher than 30%, the instrument needs a basic calibration / adjustment. | | | | | Clause  6.13 & 10.13  Manufacturer shall provide a description of the technique to determine and compensate for zero and span drift. |
| Change in zero point over maintenance interval |  |  |  |  |  |  |
| TOC |  |  | 1.5 |  |  | <3.0% |
| Change in reference point over maintenance interval |  |  |  |  |  |  |
| TOC |  |  | 1.7 |  |  | <3.0% |
| Availability |  |  |  |  | >98% | >95% |
| Reproducibility |  |  |  |  |  |  |
| TOC |  |  |  | 2.6% |  | <3.3% |

Note 1: Ambient temperature data was achieved at the range of 0 to 40 mg/m3 (5ºC to 45ºC) and at 0 to 15 mg/m3 (5ºC to 40ºC).

Note 2: Additional ranges for cross sensitivity to O2 (21 %vol. & 3 %vol.), NO2 (100 mg/m3) and HCl (200 mg/m3) were tested separately. Cross sensitivity to CH4 has not been tested.

# Description

Thermo-FID is an extractive gas analyser for the measurement of Total Organic Hydrocarbons (TOC). It uses a Flames Ionisation Detector (FID) for its measurement. The Thermo-FID represents a modular designed instrument using identical plug-in assemblies for easy service. Thermo-FID is available in different versions:

* Thermo-FID PT (Portable Thermo-FID)
* Thermo-FID FE (Thermo-FID model Field Housing),
* Thermo-FID ES (Thermo-FID model 19” Rack Mount),
* Thermo-FID TG (Thermo-FID model Desk Top)
* Thermo-FID MK (Thermo-FID model Close Coupled Probe On Stack)
* Thermo-FID KA (Thermo-FID model Cassette Mount)

All versions of the analyser have the same detector unit, gas piping system and electronics. However the Model PT uses a membrane pump instead of an injector pump. Apart from this the main differences between the models are the different designs of housings.

For further details on higher ranges available please consult the manufacturer.

# General notes

1. This certificate is based upon the equipment tested. The manufacturer is responsible for ensuring that on-going production complies with the standard(s) and performance criteria defined in this certificate. The manufacturer is required to maintain an approved quality management system controlling the manufacture of the certified product. Both the product and the quality management system shall be subject to regular surveillance according to ‘Regulations applicable to the holders of Sira certificates’.
2. The design of the product certified is held and maintained by TÜV Rheinland for certificate No. Sira MC050062/05
3. If a certified product is found not to comply, Sira should be notified immediately at the address shown on this certificate.
4. The certification marks that can be applied to the product or used in publicity material are defined in ‘Regulations applicable to the holders of Sira certificates’.
5. This document remains the property of Sira and shall be returned if requested by Sira.